UNIVERSITETET I BERGEN

# Matematisk-naturvitenskapelig embetseksamen

Subject: GFJ 372 Seismic Instrumentation

Monday, May 27, 2002,  0900-1400 hrs.

Permitted remedy: Calculator, English dictionary

A total of 100 points is possible for this exam. The number of points for each question is indicated.

# Question 1, 8 points

Explain the general principle behind the forced balanced accelerometer.

Question 2, 5 points

- What generates seismic background noise ?

- What is a typical noise displacement at 1 Hz in a quiet area?

- At which frequency is the noise amplitude at its maximum ?

Question 3, 8 points

Explain the general principle behind the flash analog to digital converter.

Question 4, 12 points

A weight of 10 g is lifted off a seismometer. This gives a pulse of 20 mV amplitude out of the signal coil. A DC pulse of 20 mA into the calibration coil gives an amplitude of 30 mV.

- Calculate the seismometer calibration coil motor constant. -- If the seismometer mass is 2 kg, the generator constant is 200 V/(m/s) and a sine wave of 10 mA (p-p) at 10 Hz is input to the calibration coil, what is the equivalent earth displacement (p-p) at 10 Hz.

Question 5, 5 points

A seismometer has a natural period of 5 s, a generator constant of 200 V/(m/s).

- How large must the earth velocity (p-p) at 10 Hz be in order to be able to get an output of 10 mV (p-p) ?

- What is the corresponding earth displacement and acceleration ?

Question 6, 3 points

A digital voltmeter has a range from -20000 mV to 20000 Mv.

- How many bits does the A/D converter in the voltmeter have.

Question 7, 10 points

Give a general description of a triggered data acquisition system like SEISLOG or GBV316. In particular, describe the following:

- Explain how the standard STA/LTA trigger works.

- Explain what effect a signal DC component can affect the operation of the trigger and how it is corrected for.

Question 8, 8 points

- What is an anti alias filter ?

- If the analog to digital converter has a dynamic range of 60 dB, what must the damping of the filter be at the Nyquist frequency ?

- If we assume a 6 pole filter and a sampling rate is 100 Hz, what is the corner frequency of the 6 pole filter ?

Question 9, 8 points

- Make a drawing of the typical elements of a modern digital seismic stations and show the interconnection of the elements.

Question 10, 3 points

A seismometer has a critical damping resistance of 5 kohm and an internal resistance of 3 kohm.

- If the desired damping is 0.6, calculate the damping resister to use.

Question 11, 7 points

- What defines a broad band sensor as compared to a long period or short period sensor ?

- Describe how to install a broad band sensor in order to get broad band performance.

Question 12, 2 points

- An amplifier has a gain of 66 dB, how many times is the amplification ?

Question 13, 3 points

A 1g accelerometer with a max output of 10 V is inclined 10 degrees from vertical.

- What is the output from the vertical channel ?

Question 14, 3 points

A digital seismic station has 20 channels of data. The sample rate is 200 samples/sec and a 16 bit A/D converter is used. The data is compressed by a factor of 2. The recording is continuous.

- How much data (Mbytes) are recorded in 24 hours ?

Question 15, 2 points

- When increasing the damping of a seismometer, how is the free period changing ?

Question 16, 2 points

- A car battery has an unloaded voltage of 12.44 V. What is the approximate % charge ?

Question 17, 3 points

- What is a typical bandwidth and dynamic range of an accelerometer.

Question 18, 3 points

- What is cross talk and in which unit is it measured ?

- What is a typical value for an analog to digital converter ?

Question 19, 3 points

- How is the self-noise of an analog to digital converter measured ?

- What is a typical noise level (counts) ?

Question 20, 2 points

- Write the poles and zero expression for the frequency response function for a standard velocity transducer.