UNIVERSITETET I BERGEN

# Matematisk-naturvitenskapelig embetseksamen

Subject: GFJ 375 Seismic Instrumentation

Monday, June 11, 2007, 0900-1300 hrs.

Permitted remedy: Calculator, English dictionary

A total of 100 points is possible for this exam. The number of points for each question is indicated.

# Question 1, 10 points

In the circuit below, all resisters have a resistance of 1 kohm.

- Calculate the current through each resistor

- Calculate the total current delivered by the power source

Question 2, 10 points

A weight of 20 g is lifted off a seismometer. This gives a pulse of 20 mV amplitude out of the signal coil. A DC pulse of 20 mA into the calibration coil gives an amplitude of 30 mV.

- Calculate the seismometer calibration coil motor constant. If the seismometer mass is 5 kg and a sine wave of 10 mA (p-p) at 5 Hz is input to the calibration coil, what is the equivalent earth displacement (p-p) at 5 Hz.

Question 3, 10 points

You have an unknown A/D converter. Set up an experiment to determine maximum voltage range and sensitivity in counts/volt. How do you calculate the dynamic range from the above two parameters.

Question 4, 10 points

- Make a drawing of the typical elements of a modern digital seismic stations and show the interconnection of the elements.

Question 5, 4 points

A digital seismic station has 10 channels of data. The sample rate is 200 samples/sec and a 12 bit A/D converter is used. The data is compressed by a factor of 2. The recording is continuous.

- How much data (Mbytes) are recorded in 24 hours ?

Question 6, 10 points

Give a general description of a triggered data acquisition system like SEISLOG or GBV316. In particular, describe the following:

- Explain how the standard STA/LTA trigger works.

- Explain what effect a signal DC component can affect the operation of the trigger and how it is corrected for.

Question 7, 8 points

- What is an anti alias filter ?

- If the analog to digital converter has a dynamic range of 60 dB, what must the damping of the filter be at the Nyquist frequency ?

Question 8, 6 points

- What is open circuit damping ?

- How do you measure it (describe qualitatively) ?

Question 9, 6 points

A 1g accelerometer with a max output of 10 V is inclined 10 degrees from vertical.

- What is the output from the vertical channel ?

Question 10, 2 points

- An amplifier has a gain of 66 dB, how many times is the amplification ?

Question 11, 3 points

- A car battery has an unloaded voltage of 12.44 V. What is the approximate % charge ?

Question 12, 3 points

A seismometer has a critical damping resistance of 10 kohm and the coil resistance is 13 kohm.

- What is the highest damping that can be obtained for this sensor and what is the damping resistor for the highest possible damping ?

Question 13, 3 points

Sketch the displacement response for a velocity transducer.

Question 14, 4 points

A seismometer is connected to a seismic recording system. How can you determine the total electrical noise (in terms of counts) generated by the seismometer and recording system ?

Question 15, 2 points

- What is a typical internal resistance of a digital voltmeter ?

Question 16, 4 points

- What is cross talk and in which unit is it measured ?

- What is a typical value for an analog to digital converter ?

Question 17, 3 points

How do we measure polarity of a seismometer ?

What is the polarity of the output voltage from the seismometer supposed to be when the ground moves down.

Question 18, 2 points

- A sine wave has 1 V p-p, what is the RMS voltage ?

- A square wave has 1 V p-p, what is the RMS voltage ?