UNIVERSITETET I BERGEN

Matematisk-naturvitenskapelig embetseksamen

Subject: GFJ 372 Seismic Instrumentation

Thursday, March 26, 1998,  0900-1400 hrs.

Permitted remedy: Calculator, English dictionary

A total of 100 points is possible for this exam. The number of points for each question is indicated.

Question 1, 5 points

A battery has a voltage of 12.5 V, a nominal capacity of 100 A-hours. If it is charged with a constant current of 0.5 A, how long will it take to fully charge the battery.

Question 2, 5 points

3 resistors are connected to a 12 V battery as shown below.

------------------

|

|

|

12 V            -------|-------

|      |      |

R1     R2     R3

|      |      |

-------------------------

R1=R2= 3 ohm, R3= 6 ohm

- Calculate the current through each resistor.

- What is the equivalent resistance to the 3 resistors in parallel.

Question 3, 10 points

A solar panel gives out 20 V. When shorted, the current is 1 A. When connected to a car battery, the voltage is 14 V.

- Calculate solar panel internal resistance.

- What is the charging current ?

Question 4, 5 points

What is the problem of ambient electrical noise pickup and how can we eliminate it ?

Question 5, 10 points

- A seismometer has a natural period of 5 secs, a generator constant of 200 V/(m/s). It is connected to an oscilloscope set to a sensitivity of 1mv pr division. How large must the earth velocity (p-p) at 10 Hz be in order to be able to see the signal on the oscilloscope with an amplitude of 6 divisions (p-p).

Question 6, 10 points

A seismometer has a critical damping resistance of 10 kohm.

- What is the total circuit resistance to get 70% of critical damping ?

The internal resistance is 4 kohm.

- How big a damping resistor must be connected to obtain 70 % of critical damping ?

Question 7, 15 points

- A weight of 10 g is lifted off a seismometer. This gives a pulse of 20 mV amplitude out of the signal coil. A DC pulse of 20 mA into the calibration coil gives and amplitude of 30 mV. Calculate the seismometer calibration coil motor constant. If the seismometer mass is 2 kg, the generator constant is 200 V/(m/s) and a sine wave of 10 mA (p-p) at 10 Hz is input to the calibration coil, what is the equivalent earth displacement (p-p) at 10 Hz.

Question 8, 6 points

- A digital voltmeter has a range from -20000 mv to 20000 mv. How many bits does the A/D converter in the voltmeter have.

Question 9, 10 points

An accelerometer with a max sensitivity of 1 g is inclined 45 degrees to the north. The sensitivity is 10 v/g.

- What is the output  on each of the 3 channels ?

Question 10, 6 points

A digital seismic station has 3 channels of data. The sample rate is 100 samples/sec and a 24 bit A/D converter is used.

The recording is continuous.

- How much data (Mbytes) are recorded in 24 hours ?

Question 11, 6 points

At about what frequency is the seismic background noise at its maximum ?

What is a the typical background noise level in displacement at a low noise seismic station at 1 Hz ?

Question 12, 3 points

How do we measure polarity of a seismometer ?

What is the polarity of the output voltage from the seismometer supposed to be when the ground moves down.

Question 13, 2 points

An amplifier has a gain of 66 db, how many times is the amplification ?

Question 14, 2 points

- When increasing the damping of a seismometer, how is the free period changing ?

Question 15, 2 points

- What is the relation between RMS and peak to peak voltage ?

Question 16, 3 points

- What is a low pass filter ?

- What is a band pass filter ?