GEOF 375

Seismic instrumentation

Jens Havskov

Department of Earth Science

University of Bergen

Norway

jens@geo.uib.no

January 2008

CONTENT

Introduction

Experiment 1  : 	Use of multimeter, DC voltage, currents and resistance

Experiment 2  : 	Power sources

Experiment 3  : 	Measurement of signals: Oscilloscope, and function generator

Experiment 4  : 	Capacitors, filters and frequency response

Experiment 5  : 	Recording analog and digital signals

Experiment 6  : 	Seismic sensors, basic tests

Experiment 7  : 	Seismic sensors, calibration

Experiment 8  : 	Digital recorders using triggered recording

Experiment 9  : 	Seismic field instruments

Experiment 10 : 	Using SEISLOG for sensor testing and seismic noise

measurements

Pensum

-----------------------------------------------------------------------------------------------------------------

Background material:

- Book: Instrumentation in earthquake seismology by Havskov and Alguacil, Springer Verlag

Introduction

The purpose of the course is to get a practical understanding of installing, calibrating, and operating seismic stations. The course consists of a series of lectures and 10 practical exercises, each resulting in a report. The main time in the course is spent on the practical experiments.

The course starts with an introduction to basic electronics, since it is assumed that the participants do not have much prior experience in electronics. The first experiments thus deals with common components such as batteries, resisters and capacitors. At the same time basic measuring instruments are used: Multimeter, oscilloscope, power supply, function generator and frequency counter.

The first experiments related to seismology (and other fields also) deals with power supplies: Batteries and solar cells. This is often an neglected part of seismic instruments.

Common to other fields also are measurement of AC signals and use of filters. In this section both analog and digital recorders are used.

With the above background, seismic sensors of all types introduced: Geophones, short and long period seismometers, broad band sensors and accelerometers. The experiments deal with all aspects of test and calibrations.

The participants are now ready to deal with complete station since all elements have been dealt with: Sensor, amplifier and recorder. Test are made with both analog and digital field recorders and a local network is examined. Time and weather permitting, there will be a visit to a field station.

The final part of the course deals with more advanced aspects of digital recording and data processing. Signals are digitally recorded from different sensors, calibration function are made and, in the computer, the signals are filtered and corrected for instrument response.  The complete background noise spectra are calculated digitally and compared to world wide measurements.

In addition to these experiments, there will also be lectures on timing systems, transmissions of signals and installation of seismic stations.

All references to chapters or sections in the following refer to the book mentioned above.

Experiment 1

Use of multimeter, DC voltage, currents and resistance

Background text: Appendix in book

The purpose is to get to know and use the most common electronic instruments.

The theoretical background is Ohms law:

Voltage = current * resistance

Multimeter

The general purpose multimeter can measure V (AC and DC), ohms, current (AC and DC). A few more advanced models will also measure frequency, db, and peak values. All of these functions will be tested.

Experiment 1.1

Voltage DC

- Measure the voltage of a 1.5 v battery to all available digits (Note: This depends on range used, use range which gives highest accuracy).

- Measure the range of voltages the power supply can deliver by varying the output voltage. Use measurements from both the built in analog voltmeter and an externally connected digital multimeter.

- Compare the measurements with analog and digital voltmeters, how accurate is the analog multimeter in comparison with the digital multimeter ?

- Does the power supply give and output voltage in accordance with specifications ?

Make measurements of the same voltage source at the same time with 5 multimeters by putting the multimeters in parallel (if possible the same model). Calculate average voltage and standard deviation and compare to manual.

- How accurate are the measurements in % (calculated from standard deviation).

Experiment 1.2

Voltage AC

-Measure the AC voltage of wall socket, compare to specifications (should be 220 V).

Experiment 1.3

Resistance

-Measure the resistance of 5 resistors. Are values within given accuracy ? Show calculation.

Experiment 1.4

Current and internal resistance

All power sources have an internal resistance. This means that when current is flowing through the power source (e.g. a battery), there will be a voltage drop internally in the power source. The voltage available to the user is therefore smaller than the unloaded voltage (or nominal voltage).

- Measure the voltage for a 1.5 V battery (unloaded before connecting it), use maximum resolution of the voltmeter.

- Connect the 1.5 V batty in series with a 10 ohm resister and a multimeter measuring current (see figure below).

- Measure the current with the multimeter.

- Using a second multimeter, measure the voltage over the resistor, the multimeter in the circuit and the battery.

- Are the voltage measurements consistent ?

- Calculate the internal resistance of the multimeter and the battery, can any of them be ignored ?

- Explain all calculations.

- The resistor is of 1/4 w type, can it sustain the current ? (power = current * voltage).

- What is the maximum current it can sustain ?

- What is the maximum current the multimeter can measure, the smallest ?

Experiment 2

Power sources

Background text: Chapter 7.

This exercise will deal with the most common power sources used: Batteries, AC power supplies and solar panels.

Battery capacity is measured in ampere-hours, meaning number of hours it will deliver 1 amp. If a higher current is used, the corresponding discharge time become correspondingly smaller. This process in not linear and a higher discharge results in a lower battery capacity. The capacity is also temperature dependent.

The capacity of a battery can be measured by the maximum charge it can give. A standard 1.5 V battery will deliver about 3-5 A when short-circuited  if in good charge. Less than about 1 A  and it has a low charge.

Experiment 2.1

State of charge of a 1.5 V battery

- Measure voltage of the battery.

- Test the 1.5 batty by shorting it with the multimeter measuring ampere DC (MAKE SURE YOU USE THE RIGHT RANGE DC 10A or 20 A), IF NOT YOU BLOW THE FUSE). Calculate the internal resistance and compare to the value obtained in experiment 1.4. Make sure you use the same battery.

- How well is battery charged ? Show calculation.

Experiment 2.2

State of charge of a lead acid car battery

The capacity of a car battery  (lead-acid) can also be measured by shorting it, however a special high capacity ammeter is required. A simpler way of measuring the state of charge is by measuring the voltage of the battery which is linearly proportional to the charge under the condition that the battery has not been charged or discharged within the preceding 5 hours. At 20 deg., a fully charged battery will have a voltage of 12.90 V and at zero charge it is at 11.80 V.

- Determine the state of charge in % for the car battery. Assuming the battery has a capacity of 60 amp-hours, for how many hours will it be able to operate a light bulb with a current drain of 0.2 amps ?.

The state of charge of a lead acid battery can also be determined by measuring the density of the acid. Use the density measuring devise on the car battery, compare results with above. BE CAREFUL NOT TO SPILL ACID.

Experiment 2.3

Charging  batteries

A lead-acid battery should not be discharged completely since this reduces the life of the battery. A normal lead-acid battery will only sustain a few complete discharges.  Charging the battery is not a linear process. Initially, the battery will be charged with a high current (e.g. 5 A), however, when the charge increases, the charging current decreases to about 1 A although the battery is far from charged. From then on it will take at least one day to fully charge the battery. To completely charge the battery (seldom done), it must be charged for some hours at a higher voltage (15.5 V). Some battery chargers automatically do this.

Some  batteries are almost completely sealed and can sustain more charging cycles than a car battery and therefore used in many instruments and at places where it is inconvenient to use an open lead-acid battery. However, the charging might be different from the lead-acid batteries, since not all are lead-acid type.

Additional gel cell lead acid batteries are available. Connect a 2 ohm power resistor to one or two of these and measure the voltage with and without the resistor. Note the voltages.

- Determine the internal resistance.

- What does the values tell you about the state of charge ?

Select a battery that is not fully charged. Connect a power supply at 13.8 V and measure the charging current and voltage.  MAKE SURE TO USE CORRECT POLARITY SINCE  OTHERWISE THE POWER SUPPLY MIGHT BE DAMAGED.

- What is the charging current and what does it tell you ?.

- Leave the battery charging. Repeat the measurement after at least 15 min, what is the difference in charging current and why ?

Experiment 2.4

Power supplies

The power supply converts 220V AC to e.g. 12V DC. It will have an internal regulator so that the voltage does not drop when the load changes unless loaded to a higher current that it can sustain. All power supplies will have some AC left (ripple, see Appendix) which can disturb the instruments. Normally this should be very small and cause no problem.

- -Connect a 10 ohm power resistor to the power supply. Set the voltage so high that the load of the power supply is at least 50 % of maximum load in ampere (indicated on power supply). Measure the AC voltage (ripple) coming out of the DC power supply, with and without load, and compare to specifications in power supply manual.

- Measure the current at the same time with the meter on the power supply.

- What happens if you short the power supply, use both 15 and 30 V ranges ( measure current drawn and voltage out) ?

- The power supply has current limitation. Explain what its purpose is.

- Connect a 10 ohm power resistor to power supply, set voltage to 10 V.

- What is the current, when current limiter is set to max ?

- Set the current limiter  to 0.5 A by shorting the supply and measure output on ampere meter until it is 0.5 A. Reconnect to  10 ohm resistor.

- -What is the output voltage ? Explain observations.

- The power supply now gets hot, why ?

Experiment 2.5

Solar panels

Solar panels generate a voltage which is proportional to the light input. They must therefore have a regulating circuit when charging batteries to avoid overcharging in strong light and draining the battery through the solar panel in low light.

- Measure output (V) of solar cell turned towards light and turned away from light (in lab). Find maximum current it can deliver in the lab and calculate internal resistance for all cases. What is the maximum power it can deliver in the lab. Can it be used to charge a 12 V battery ? If so, what is the charging current.

- Repeat measurements outside, towards and away from sun, describe if full sun or no sun !!! - Calculate charging current at 14 V when placed outside (remember to take internal resistance into account).

- How is the internal resistance dependent on the light intensity ?

- Measure the area of the solar panel and calculate the efficiency of the panel in full sun. Compare to what is expected (book section 7.8.1).

- -If a battery self discharge completely in 2 years, what is the self discharge current ?

- Can the solar panel used inside keep this battery charged ?

Experiment 3

Measurement of signals: Oscilloscope and function generator

Background text: Appendix in book

This section will deal with signals, how to generate and measure them.

The function generator generates oscillating electrical signals of various frequency and amplitude. Generally sinusoidal signals are used, but most generators can also produce square waves and triangular waves. The signals are used for testing electronic circuits and in this course, specifically to generate test signals simulating the earthquake generated signals. The generators usually have a frequency band from a little below 1 Hz to some MHz. When used in connection with seismology, it is important to use generators which can generate low enough frequencies, since many standard instruments will not go below 1 Hz. For use in seismology, the instrument must be able to go to at least 0.1 Hz and preferable lower.

The oscilloscope (OC) is used to graphically visualize an oscillating electric signal (Volt AC) and can be used to measure both the amplitude and frequency of the signals. However, it can also measure DC signals. Like the multimeter, the oscilloscope can also be set to different sensitivities. Most older types oscilloscopes are analog, while modern oscilloscopes are digital.

For details of operation of these instruments, look in instrument manuals.

NOTE: If connecting equipment with single cables instead of a coaxial cable (like a generator to an oscilloscope), make sure ground is connected to ground. In some cases instruments have ground connected to AC ground and if that is the case, the signal might be shorted if ground is not to ground.

Experiment 3.1

Measure signals with an oscilloscope

- Set the sensitivity of the OC to 1V/cm, the sweep time to 0.1 sec, 1 sec, 10 s or whatever is the largest, note the settings (speed and sweep time) time used. Time settings are in seconds/unit while sweep times is total time across screen.  What is observed ?

- What is the range of voltages and sweep times that can be used and what is the approximate frequency range which can be measured with the resolution of at least one division per period and not less than one whole period pr screen.

Experiment 3.2

DC and the oscilloscope

- With a sweep time of 1 sec, connect the 1.5 V battery, what is observed ?

- Measure the battery voltage, try the AC and DC switch of the instrument and explain what is does.

- Invert the battery poles and observe what happens.

Experiment 3.3

Measuring AC with the oscilloscope

- Connect the OC test probes to the function generator set at 10 Hz and 10000 Hz, any voltage.

- Adjust gain and sweep time and measure the voltage and frequency, how do the values compare to what is set on the generator ( voltage not always shown since not all generators have display). Make sure you indicate the accuracy of the measurements.

Experiment 3.4

RMS voltage

- The multimeter measures AC RMS values (RMS is root mean square). Connect the function generator (FG) to the multimeter, adjust the frequency to 100 Hz and the amplitude to 1V RMS on the multimeter. Connect the same signal to the OC (DC coupled) and measure the frequency and amplitude, explain the difference.

- Repeat the experiment with square waves and triangular waves. For triangular waves it can be shown that theoretically, Vrms = V / √3.

- What do you expect the relation to be for square waves and why ?

- Now explain differences in above measurements.

- Does the multimeter measure RMS accurately for triangular waves, square waves and sine waves ?

- Repeat the experiment with a 1.5 V battery in series with the signal, what do you observe, explain the difference. The FG might have a DC level adjustment, try it, which range does it have, explain its use.

Experiment 3.5

Noise pickup

- Any cable not grounded will pick up 50 Hz radiated from the power cables nearby. This is a source of noise in electrical circuits. Connect a 1m long unshielded cable to the + input of the oscilloscope and multimeter + at the same time and measure the amplitude of the 50 Hz signal picked up.  Measure with and without the two ground connections together. Why is there a difference (see also 3.4).

Experiment 4  Capacitors, filters and frequency response

Background text: Chapter 6 and appendix in book, RESP program in SEISAN manual.

Any electrical circuit which has an input and an output (like multimeter, oscilloscope, amplifier or a combination of electronic elements such as resistors and capacitors) might modify an incoming AC signal in both amplitude and phase. The relation between the output and input signals is called the frequency response of the circuit. We assume that we are dealing with linear systems meaning that if e.g. we double the input signal, the output will also double. In this exercise we will look at the frequency response of both the voltmeter and the oscilloscope and also construct a simple filter.

A seismometer can be considered in a similar way except that the input is ground motion and the output is voltage, we will look at that in a later experiment.

The RESP program calculate response functions for seismic systems with both sensors, filters etc. In this experiment, we will only use the filter part of RESP.

Experiment 4.1

Frequency response of multimeter and oscilloscope

- Connect the FG to the multimeter, adjust the frequency to 100 Hz and the voltage to 1 VRMS. Now measure the amplitudes at the following frequencies: 80, 50, 40, 25, 10, 5, 2, 1, 0.5, 0.2, 0.1.

- -Plot the voltages as a function of frequency on a log-log plot. What does the curve show, what does it tell about the use of a multimeter for measuring low frequency AC signals ?

- For which frequencies are the measurements reliable.

- In DC mode, it is possible, using a multimeter, to follow the voltage of a slowly varying signal, however when the frequency increases, the averaging of the multimeter will make this measurement inaccurate. Determine this frequency by determining peak amplitudes and comparing these to the know input amplitude. Explain results.

- Give definition of dB scale.

- Repeat the experiment with the OC in AC mode and DC mode, what is observed ?

Experiment 4.2

Time constant of a capacitor

- Connect a large capacitor ( ~  50 micro F, note the capacitor might be polarized so it has a + and a - side, do it right !!) in parallel with a large resistor ( ~ 1 Mohm). Connect to a power supply (10 V) and a voltmeter in parallel so the voltage can be monitored without too much load. Disconnect the power supply and determine the time it takes to reduce the voltage to 3.7 V. This is the time constant (see Appendix). Calculate the time constant from the capacitance and resistance.

- How does the measured and calculated values compare ?

- Can you theoretically disregard the internal resistance in the voltmeter ?

Experiment 4.3

Filter

- Construct a low pass filter with a cutoff frequency larger then 20 Hz , show a drawing of the filter with input and output. Make sure you do not used a polarized capacitor, why ?

- What is your calculated cut -off frequency (- 3dB point) ?

- Connect a signal generator to the input and one trace of an OC to the input and one trace of the OC to the output and vary the frequency. Now measure the amplitude and phase shift of the output signal relative to the input signal. Plot the measured values of gain and phase shift as a function of frequency and also plot the theoretical curve on the same graph.

- How do they compare ? The theoretical curve can be calculated with RESP or calculated manually.

Note on phase shift: A sine wave can be described as

sin(2*π*f)

where f is the frequency. If the phase shift is s, the expression is

sin(2*π*f+s)

Comparing the two signals on an oscilloscope then means that if s is positive, the output signal is shifted backwards relative to the input signal.

Experiment 5

Recording analog and digital signals

Background text: Chapter 4 and 5, SEISLOG manual.

Seismic signals converted to voltages and other signals are generally recorded continuously for analysis. So far in the experiments, we have only dealt with signals measured instantaneously. Modern digital oscilloscopes can store some limited time sequence and even on the traditional oscilloscope, it is possible to see some time sequence (afterglow). However for further measurements, it will often be desirable to be able to see seconds or minutes of data plotted, either on paper or on a screen. So in this exercise we will jump to the final recording of signals although we have not yet dealt with sensors and amplifiers, since it is simply practical to be able to use some of the recording tools we have available.

For analog recording, we have a heliocorder. The heliocorder can record signals for 24 to 72 hours on one sheet of paper. The heliocorder is one of the most common analog recording devices in seismology, although used less now.

All modern systems record digitally. In these experiments we will use the SEISLOG and SEISMOWIN recording systems. SEISLOG is capable of many things to be dealt with later, however in this experiment we will just use the ability to record and plot a signal.

Experiment 5.1

Heliocorder

Connect a signal generator to the heliocorder and measure the frequency response with the gain set to 0 dB. Use the frequencies 0.1, 1.0, 5.0 20.0 40.0 and 100 Hz. Describe the frequency response and calculate the sensitivity in m/v.

Experiment 5.2

SEISLOG A/D converter sensitivity

The SEISLOG system used in this exercise has an external A/D converter connected to the PC via a serial line. Different digitizers might be used. Familiarize yourself with the SEISLOG operation, including set up of digitizer and recording (see manual). At this time, do not record in ring buffers or set up trigger, this will be done in a later exercise. Test the main functions under view, configure and trig and make sure you understand  most parameters on the main display. It is also assumed that you are familiar with main functions of SEISAN to be used for inspecting the data.

For 2 different A/D converters, do the following:

- Short input to get DC shift.

- Connect a variable voltage power supply at 0.2 V to channel 1 of the digitizer.

- Measure the number of counts using monitor screen, possibly correct for DC shift.

- Calculate the A/D converter sensitivity in counts/V. Also calculate how many volts one count represent (LSB: Least significant bit).

- Increase the voltage slowly from about 0.2 V until the numbers (counts) become constant. The corresponding voltage is the maximum the converter can use. Note down maximum counts and corresponding voltage.

- Calculate the number of bits for the converters.

Experiment 5.3

Recording a signal with SEISLOG and determine antialias cutoff

Connect a signal generator to channel 1, frequency 10 Hz, unknown amplitude (to be determined by the recording). The signal should be visible using the View facility.

- Start and stop recording with Trig/Simulate event facility (record in WOR).

- Plot the signal with SEISAN.

- Find the maximum amplitude (peak to peak) in counts and convert to V using results from 5.2 (if same digitizer, else get value from a manual).

- Measure the input amplitude and frequency with an oscilloscope, is it correct ?

- Measure the recorded amplitude at frequencies e.g.10, 15, 20, 25, 30 and 50 Hz, either by using the monitor values (use max amplitude) or by recording the signal in a file and checking it using SEISAN.  In this way the frequency response curve can be determined. NOTE: The frequencies used will depend on the sample rate used so some experimentation is needed.

- Plot the amplitude as a function of frequency and determine the cutoff frequency of the antialias filter. Note, some digitizers do not have antialias filters.

- Calculate the antialias filter damping at the Nyquist frequency, is it ok ?

- Check the sample rate by zooming.

Experiment 5.4

A/D converter internal noise and cross-talk

Most converters have an internal noise, which will limit the dynamic range. The purpose is to measure the true dynamic range. There is always some overplay from one channel to another which is also called cross-talk. The purpose is to measure both quantities.

- Short the input channels and make a recording. The signal now seen represents the internal noise. Measure the average AC amplitude signal (by plotting) and calculate the converter dynamic range. Compare to manufacturer specifications.

- Connect a signal generator to channel 1, a resistor of ca 5 kohm to channel 2 and short channel 3. Put in a 10 Hz signal at close to clipping level and measure the 10 Hz signal amplitude in channel 2 and 3. Calculate the percentage overplay to the other channels, convert to dB and compare to specifications.

Experiment 6

Seismic sensors, basic tests

Background text: Chapter 2 and 10

The purpose of this exercise is to become familiar with the most common type of seismic sensors and be able to make simple checks.

Experiment 6.1

Seismometer clamping

The most common short period seismometers are Ranger SS1, Teledyne S13 and Mark Products L4. In this exercise, all 3 must be used as well as at least one high frequency geophone and one more seismometer. Many seismometers have a clamping device used when transporting the instrument.

- Explain the use of the clamping device and determine which of the units you test have a clamping device. If there is no clamping device, how can the seismometers safely be transported ?

Experiment 6.2

Horizontal and vertical seismometers

Seismometers are normally horizontal or vertical. Which of the seismometers can operate in both ways and what must be done to change between vertical and horizontal operation ? Try it with the model, which can be changed in the field and set up the seismometer.

Experiment 6.3

Basic parameters

- -Measure basic parameters of the 5 test seismometers (Ranger SS1, Teledyne S13, L4C, LP and a geophone): Free period, resistance of signal coil, resistance of calibration coil (if present). NOTE which seismometer serial number is used. The signal coil will normally have a much larger resistance than the calibration coil.

- When measuring the resistance of the signal coil, it is recommended to clamp the seismometer mass, why ?

- What is observed if the mass is not clamped, explain.

- Describe how you measure free period (by counting swings or measuring with an OC or a recording program). Note that the L4C seismometers might have a damping resistor built in !!!. The signal coil resistance should be 5500 ohms, if not it has a built in resistor. Determine it by assuming that the measured resistance is the parallel resistance of the coil and damping resistors.

Using the S13 or Ranger seismometer, connect a multimeter in DC mode (most sensitive range).

- How much do you have to move the sensor to get an output (measure in multimeter DC in mv range)?

- How much output (volts p-p) do you get by tilting the instrument slightly (less than needed for the mass hitting the stop) ?

- Redo with a geophone, can you estimate ratio of generator constants ?

Experiment 6.4

Long period seismometer

For the long period seismometers (vertical and horizontal), explain how the period is adjusted. Try to adjust the period of the horizontal to the longest possible value, how low is it possible to get ?

Experiment 6.5

Polarity

In order to make fault plane solutions and modeling of seismic signals, it is important that the polarity is ok. The convention is that an earth movement UP, to the EAST  and NORTH shall give positive signals and finally result in a positive signal on the display (screen, plot or seismogram). The first step is to identify the positive terminal on the seismometer. This can be done with a multimeter, or better an oscilloscope while doing a tilt test. SEISLOG or SEISMOWIN can also be used. This is done by rapidly tilting the seismometer, which means that the seismometer moves UP when using a vertical seismometer or horizontally in direction of rotation for a horizontal seismometer. Notice the first pulse on the measuring device and determine the positive terminal on at least 3 different seismometers of which one is a horizontal.

Experiment 6.6

Broad band sensor, optional, sensor might not be available.

The broadband band sensor can be understood as a usual velocity sensor with an extended frequency range in the low frequency end. The long period sensor is usually going down to about a period of 20 seconds, while broad band sensors can go down to a period of several hundred seconds by using sophisticated electronics. In this experiment, a Guralp broad band sensor will be used.

- Check which sensor is used (CMG40T or CMG3T) and get free period and  generator constant from the notes or the manual. How are the values compared to a Ranger seismometer and the long period seismometer ? Which is the most sensitive ?

- Connect the broad band sensor to the control box and check the mass position (explain what mass position is). If the mass position is out of spec (see manual), re-center the mass. How is it done ?

- Check polarity as in 6.5 using  the meter in the control box. Estimate period by tilting and measuring the period by counting seconds for a half period.

- Connect the sensor to the SEISLOG system and monitor the channels with the View facility. What is the dominant period of the signal ? Cpmåare to the broad band signal at the BER station.

- Connect a voltmeter to the output of the sensor. Can you follow the signal variations with the voltmeter ? If you can follow the signal, what is the dominant period ?

Experiment 7

Seismic sensors, calibration

Background text: Chapter 2 and 10

This exercise is a continuation of exercise 6 where more advance features are tested. In particular, the procedures of calibration are tested and the accelerometer is also used.

Experiment 7.1

Calibration coil

The calibration coil (if fitted) is used to generate step functions simulating a step earth acceleration. Connect the oscilloscope (or recorder) to the signal coil and determine a "suitable current" to send through the calibration coil to get a "suitable output" (meaning that calibration signal is well above the seismic background noise) on the oscilloscope. Note the values. You can always use a 1.5 V battery. BE CAREFUL, IF USING A POWER SUPPLY, TO START WITH A LOW CURRENT (<100 mA) so you do not burn off the calibration coil.

- What sweep time is best to use ?

- What is the highest current you theoretically can put through the calibration coil with the 1.5 V battery ?

Experiment 7.2

Determine motor constant of calibration coil

- Determine the motor constant of the calibration coil for the S13 seismometer (vertical or horizontal) following instruction in section 10.2.3. Compare to specifications.

Experiment 7.3

Determine seismometer damping

The damping is an important parameter which must be set right in order to record the signals without too much over swing. The damping is normally set to (0.7 * critical damping). Use a seismometer and a geophone.

- Connect the seismometer to an OC or SEISLOG. Make the seismometer swing by either giving it a small kick or sending in a calibration pulse (if calibration coil). Determine two subsequent amplitudes and calculate open circuit damping (section 10.2.2).

- Explain what open circuit damping is, can it be ignored for the seismometers tested ? Repeat experiment with a geophone.

- Repeat the experiment (seismometer only) with a 'suitable resistor' (1-30 kohm depending on sensor) in parallel with the seismometer signal coil so the over-swing is about 20-30 % (meaning that following amplitude is 20-30% of previous amplitude). Determine the ratio of two subsequent swings and calculate (section 10.2.2) critical damping resistance (remember open circuit damping). If it is difficult to read the signal on the OC, use analog or digital recording. Compare to specifications.

- -Calculate the external resistance needed for critical damping and for 0.7 of critical damping.

- Use a resistor higher than needed for critical damping and measure the free period. How does it compare to the free period measured without an external resistor. Compare to the theory (section 2.7).

When a damping resistance is connected to a seismometer, the output from the seismometer is reduced. This can be specified by giving the seismometer loaded generator constant which is the generator constant as seen from outside when the damping resistor has been connected. Below is an example of the calculation.

- A seismometer has a generator constant of 300 V/(m/s), an internal resistance of 5000 ohms and a damping resister of 3000 ohms, what is the loaded generator constant ?

NB: If you do this experiment with a digitizer, make sure you take into account the input resistance of the digitizer, which might have a damping resister built in (check manual). In order to eliminate the effect of the digitizer input resistance, connect the sensor to the digitizer in series with a large resistor (more then 100 k).

Experiment 7.4

Seismometer generator constant

The seismometer generator constant can be calculated if all seismometer parameters are known. For an unknown seismometer, it can also be calculated if only the inertial mass and critical damping resistance have been measured (section 2.4.1). Estimate the inertial mass of one of the seismometers in 7.3 and determine generator constant by calculation.

The generator constant can also be measured directly by placing the seismometer on a shaking table and measuring the output for a given shaking table velocity.  We have no shaking table but it is possible to compare two seismometers by simply comparing the voltage outputs generated with the earths background noise, the earth is then the shaking table.

Connect two different seismometers to two channels of a high gain digitizer. Place the seismometers close to each other. The two signals should be very similar in shape if the seismometer period is the same, otherwise thy might look quite different. Record the signals. Display the signals filtered in a frequency band above the natural frequency of  both sensors.

- Why filter ?

- How similar are the signals with or without filter ?

- What is the ratio between the two amplitude ?

- What is the ratio of the two generator constants ?

- Compare to specifications.

Experiment 7.5

Frequency response of seismometer

Use a SP seismometer where all constants are known and connect a damping resistor (corresponding to a damping of 0.7) to the signal coil. Connect the calibration coil to the frequency generator. Set the frequency to 1.0 Hz, connect the oscilloscope (or recorder) to the signal coil and slowly increase the amplitude until a clear signal is seen. Now connect the oscilloscope to the output of the generator and measure the voltage drop from which the calibration current can be measured (section 10.2.5).

- Why cannot this be measured with the multimeter (at least below 10-15 Hz) ?

- Reconnect the oscilloscope to the signal coil. Assuming that the generator gives the same output voltage (=same calibration current) at all frequencies, measure the output of the seismometer at frequencies  0.2, 0.5, 1.0, 2.0, 4.0, 8.0,  15.0, 20.0 and 30.0 Hz. At each frequency, determine the earth equivalent displacement (section10.2.5). Using the measured output voltage, calculate for each frequency the gain of the seismometer in V/m.

- Using the theoretical values, calculate fraction of critical damping, and the loaded generator constant (see 7.3) and calculate the theoretical response with the RESP program (see SEISAN manual).

- Compare the theoretical values with the measured values, how do they compare ?

Experiment 7.6

Accelerometer testing

The accelerometer measures the earth accelerations from DC to about 50 Hz without significant change in gain. The instrument therefore can measure static changes in the gravity field as well as dynamic changes.  When the accelerometer is placed horizontally, theoretically output from all 3 sensors should be zero. This is not possible in practice due to leveling problems and DC shifts in the internal electronics.

- Using one of the 2 accelerometers, connect power and using a multimeter, determine the smallest simultaneous voltage out of all 3 sensors, compare to manual.

- The simplest way to check the static properties of an accelerometer is by making a tilt test thereby changing the static acceleration in the 3 component directions. Tilt the accelerometer towards North at an angle of 30 deg (or any angle easily set up, but smaller than the angle which will saturate the horizontal sensor), measure output of all 3 channels, tilt 30 deg to the South, East and West and measure output of all 3 sensors.

- What is the sensitivity in V/g. What is the maximum acceleration the accelerometer can measure assuming the max accelerometer output is 2.5 V (explain your calculations) ?

- How is the polarity ?

- Do the tests indicate that the accelerometer is OK ?

Experiment 8

Digital recorders using triggered recording

Background text:  Chapter 5 and SEISLOG manual.

A digital recorder can simply be a device recording a signal continuously in digital form, as it has been used in experiment 5. The problem with continuous recording in seismology is that very large amounts of data will have to be recorded. Usually there is a detection program on the recorder continuously checking the signal levels and only recording time intervals where signals where the detector recognizes a signal different from the background noise. This way of operation is the most common, specially on field recorders, where due to the requirement of low power consumption, the storage capacity might be limited.

The current development of increases in capacity of memory and hard disks, reduction of physical size and power consumption has partly resulted in more systems using continuous recording, particular for system only recording 3 channels. However, most systems will at the same time use event detectors, simply to be able to get information about at which times seismic events are likely to occur.

Experiment 8.1 (textbook exercise)

SEISLOG parameters

The trigger on the SEISLOG recording system is using several parameters.

- Explain the principle behind the trigger including the use of STA, LTA and trigger ratio. Explain the use of preevent memory, post event recording and array propagation window.

- On a system recording 3 channels, 4 byte integers and has a sample rate of 50 Hz, how many Mbytes does one day of recording represent.

- Assuming 1.5 Gbytes is allocated for continuous recording, how many days of continues recording is there room for.

Experiment 8.2

SEISLOG testing

Use the test SEISLOG system. Connect a 1 Hz seismometer (on pier) to channel one of the system. Use a high sensitivity digitizer. Start the system.

- Check the signal with command View, does it look right ? Explain !

- Set the trigger to work on only channel 1 Config/Administer channels and Trig), use no filters. With the default values, find optimal trigger ratio for triggering without being too sensitive. Jump at 2 m's distance to test.  Note down. Remember to start and stop the system when trigger parameters are changed.

The filters are used to improve signal to noise ratio for the earthquake signals which might be different from the background noise signals. Using a recorded noise signal, do the spectral analysis of the raw signal using SEISAN.

- What frequency band is dominant ?

- Set the filter to reduce the noise. Plot the filtered and unfiltered signal. How does the filter work ?

- Repeat the trigger test with filters set, do you have to change the trigger ratio or other parameters to make the system equally sensitive ?

A small local network is in operation around Bergen. It consists of 3 stations, one about 20 km to the south (EGD, 1 SP Z), 20 km to the NW (ASK, 1 component SP) and 1 sensor in the lab  (SP Z). The signals from stations ASK and EGD are transmitted in analog form using FM modulation and telephone lines.

The network is recorded on the Windows SEISLOG located on the 2. floor with the heliocorders for EGD and ASK.

Experiment 8.3

Testing triggering on station BER

- Put a new paper on the heliocorder. Connect a 1 Hz seismometer through an amplifier to the heliocorder. Make sure you turn the gain high enough to clearly see the background noise. Jump near the sensor to check that it works. Test different filter settings on the amplifier.

- Which filter gives the lowest noise while still making the sensor sensitive for jumping signals ? Use this filter for the recording.

- With which filter is it possible to see the microseismic background noise ?

- Set up a SEISLOG system with the same 1 Hz sensor, set the filter to the band most favorable as observed on the Helicorder or with noise tests. The Helicorder and Seislog must have the same filter setting. Let the recorders work for 24 hours, make sure start time is noted on the heliocorder and that the Seislog system has approximately the correct time, if a GPS is not connected. After 24 h, compare the triggers on the seismogram with the triggers on the Heliocorder.

- How many triggers are seen with SEISLOG ?

- Are the same events seen on the Helicorder or are there more or less ?

- Any explanation for deviations ?

Experiment 8.4

Local Bergen network triggering

The local Bergen network will only trigger if two out of 3 stations trigger. Check with network triggers for the same 24 hours.

- How many of the network triggers are seen with experiment 8.3 ?

- Are the any triggers with the network not seen under 8.3 ? If so, why ?

Experiment 9

Seismic field instruments

Background material: Chapter 5 and instrument manuals

Most seismic stations are permanent. However, there is often a need for recording with portable instruments. This is e.g. the case if aftershocks are recorded or the micro seismicity is to be studied in a particular region. Another important use of field instruments is for recording of controlled explosions for crustal studies.

Most instruments are digital and work more or less like the SEISLOG system. The main difference is that they have to be housed in weather proof cabinets and be designed to operate on low power. Usually they are powered by internal batteries (short time use, 1-2 days), car batteries (2-4 weeks of use) or batteries in connection with solar cells for continuous power support.

Analog systems are rarely in use despite ease of operation, the availability of simple continuous recording and the advantage that the signals are observable without processing.

Experiment 9.1

Digital GBV316 field instrument

In this experiment, the GeoSys GBV316 digital seismic recorder is used. It works in principle like the SEISLOG system, however this is a small portable unit to take in the field. It serves the same purpose as the MEQ800. The GBV316 contains 3, 4.5 Hz geophones, amplifier, digitizer clock, battery and memory. Like SEISLOG, it can also be time synchronized with an external GPS clock. The trigger system is set up like on SEISLOG with STA and LTA, however the recording is in solid state memory (flash card) so there is no continuous recording, only triggered events can be recorded. The recorder is one of the smallest of its kind on the market, and reasonably priced, which has been achieved by not using state of the art components. The A/D converter is 16 bit and the memory is limited to 2-10 Mb. However, the combination of 16 bit and 4.5 Hz geophone maximize the performance. Today the instrument is outdated due to its low memory and 16 bit digitzer, but it can still be used for smaller expricents.

The instrument is communicating with a PC through the AllView (AV) program, which is used to set up the recorder and transfer and inspect data. Once the data is transferred to the PC, it can be analyzed by AV or converted to SEISAN format for more detailed analysis. The GBV316 can also be connected to a SEISLOG system for continuous recording. The GBV316 is now getting a bit outdated.

- Read AV manual section 1

- How long can the GBV operate with and without GPS using the internal battery (it is assumed that the GPS  is running 10 % of the time).

- Connect the GBV316 to a PC with AV installed.

- Turn on the power of the GBV316

- Go to directory c:\av, type av and the program starts up (could be directory d:\av).

- Type S for single instrument

- Type L for login, password is *, you are now logged into the instruments and the main menu appears

- Check Recorder display, and select Analog signals. You will now see the count levels.

- How much does the count level change and what is the average amplitude ?

- Touch the instrument to see the levels changing, this might trigger the instrument as indicated by the yellow LED being illuminated and an orange REC will appear at the lower left hand corner on the PC screen.

- Go back to main menu with ESC

- Check trigger parameters from menu Recording settings, note down the settings, are they 'reasonable' ?

- Select Event Recording Files and check events recorded, how many events are recorded ?

- Download an event and plot it (use online help or manual)

- Check how many event are recorded by looking at the instrument display

- How much memory is used, check instrument display ?

- What is the voltage of the internal battery ?

- If the memory is 2 Mb, sample rate 50 Hz and the data compression is 1:2, how many events of 100 sec duration can be recorded ?

- Turn the instrument upside down in order to lock the vertical channel. How much noise (counts) is the digitizer making on the vertical channel ? Note, the noise is not the DC signal, but the AC noise so the variation in the counts must be observed.

Make sure the instrument is properly set up and trigger easily. Make one recording on the pier, then turn the instrument off and take upstairs. Make another recording on the 2 floor.  Download event files from both recording to the PC and convert to SEISAN format using program gsrsei (could be gsrsei1 for some models).

- What is the approximate difference in seismic background noise downstairs and up upstairs ?

- Note changes in frequency.

- Note if changes in amplitude and frequency is the same on all channels.

Experiment 9.2

The SARA SL07 recorder

This recorder is the newer development of the low cost recorder with comparable specification to the GVB but with a 24 bit digitizer and up to 4 Gb of removable memory. With experience gained from previous experiment, you should now be able to set up and check the recorder by just using the manual.

- Set up and check recorder, describe the parameters you can check.

- Set up to record in continuous mode, 10 minutes files. Record a couple of files and transfer to the PC using the removable flash memory. The recording format is GSE which can be plotted directly by SEISAN.

- Disable continuous mode and set up in trigger mode. Make the instrument trigger and transfer the files to the PC and check the data by plotting.

Experiment 9.3

The Guralp recorder

The Guralp CMG-6TD is low cost broad band sensor combined with a digitizer and a recorder using flash memory as recoding media like GBV316. The difference from  the GBV316 is a better sensor, digitizer and higher recoding capacity (up to 8 GB). In addition it uses less power. The recorder is mainly intended for continuous recording. The recorder can be used with SEISLOG in which case only digitizer and sensor is used. The recorder also has its own  software SCREAM from the manufacturer. For more details, see manual. Note, baud rate might have to be set to 9600.

Power consumption

- Measure the power consumption of the recorder, compare to specifications.

Experiment 9.5

SCREAM software

Connect  the recorder to a PC and use the SCREAM software to check the instrument.

- What is the sample rate initially set up ?

- How is the sample rate changed ?

Set up the display to show the signal graphically. Make a small  tilt of the instrument and monitor the vertical channel.

- What is the approximate period of the instrument ?

- Is the polarity correct

If the recorder has 4 Gb of memory, the sample rate is 50 Hz, compression level is 50 %, for how long can the recorder record continuously ?

Experiment 10

Using SEISLOG for sensor testing and seismic noise measurements

Background text: Chapters 3 and 6  and SEISLOG manual.

With digital recording, it is possible to manipulate the signals from one sensor so it resembles another sensor. This means that the distinction between LP and SP sensors to some degree disappear, since, when knowing the response function of one, we can convert to another limited by the signal to noise ratio. The broad band sensor can simulate both the LP and SP sensor and therefore replace both. A broad band sensor is based on the forced feedback principle and can in some cases be considered an advanced accelerometer, except it has been electronically modified to give velocity as output.

The reason for still using SP sensors is mainly cost and simplicity since in many cases (micro earthquakes) there is no need to record low frequencies (less than 0.2 Hz). In this experiment we will test how similar signals we can get out of a very cheap geophone, a traditional 1 Hz seismometer and an accelerometer.

When installing seismic stations, the quality of the recordings depend to large extent on the local background noise. It is therefore important to have a standard to compare with when measuring background noise. The most widely used standard is the Peterson curves (Chapter 3) which however is given in acceleration power spectral density, something which has to be calculated. The earlier measurements done in this course has been in nm displacement and can be converted, to be compared to the Peterson curve. In order to get the whole curve, spectral analysis must be done, which it the aim of experiment 10.3.

Chapter 3 explains how to convert from a single measurement in e.g. nm to power spectral density of acceleration. To simplify, we will only consider 2/3 octave filters, meaning that we assume our single frequency measurements have a relative bandwidth of about 0.5. The 5 to 10 Hz filter is a one octave filter and has a relative bandwidth of 0.7. Assuming a 2/3 octave filter, the approximate relation between the displacement power Pd and the rms displacement d is

Pd = d * d / f0                                                 (1)

where f0 is the filter center frequency. If we assume the filter center frequency to be equal to our measurement frequency and the rms amplitude equal to the amplitude, we can calculate the power spectral density of displacement from above equation. Since our measurements are in displacement and the Peterson curves are in acceleration, the equation to calculate the acceleration power spectral density from the displacement amplitudes is

Pa = d * d * 16 * pi**4 * f0**4 / f0                       (2)

In order to get the level in dB relative to 1 (m/s**2)**2/Hz, we take the log of Pa and multiply with 10.

It is of course also possible to convert from values on the Peterson curves to displacement using (2).

This calculation is approximate. To be more accurate, the filter width must accurately be estimated and the relationship between the amplitude and the rms amplitude be established.

Experiment 10.1

Sensor comparison

With the SEISLOG, connect a Ranger seismometer to channel 1, the 4.5 Hz geophone to channel 2 and an accelerometer (Z) at the 3. channel. Remember damping resistors. Make sure polarity is ok by jumping.

- Plot one channel at a time, make some noise to see if the sensors react.

- Make a file of about 15 s duration, jump a bit to make a pulse.

- Check with EEV if the 'event' has been recorded.

- Plot signals. Use a short enough window so signal shapes clearly can be seen.  How do the signals compare in frequency and shape.

- Plot with a band-pass filter of 10 to 15 Hz. How do the signals compare in shape now ?

- Which 2 signals are most similar ?

Experiment 10.2

Response correction

- Make response files for the 3 channels recorded, remember to put files in the CAL directory if on the local SEISLOG PC, else put calibration files in your own directory and set environmental variable LOCAL_CAL to the path to your own directory. This could e.g. be setenv LOCAL_CAL /home/users/ (unix). Make sure the accelerometer component code is A  Z (2 blanks between A and Z).

- Now plot the 3 channels again using multitrace mode. Remove the response (use option g and filters 1-5 Hz and filters 5-10 Hz) and compare the 3 channels now showing displacement.

- Explain what you have done and what you expect, give a table with all calibration values used?

- How different are they, any difference using the two frequency bands ?

- What is most comparable, the noise or the 'event' ?

- Make the noise spectra (N option) using traces for each sensor, use only noise section. In which frequency bands do the spectra agree ?

- What does it tell you about the frequency bands in which the instruments can resolve earth noise ?

Experiment 10.3

Noise measurements

Connect a broad band sensor to Seislog using an external digitizer or a broad band sensor with a built in digitizer. Check with View that it works, 2-3 second background noise should be clearly seen. Now record the background noise two minutes. DO NOT MAKE 'LOCAL EVENTS'. Make a response file for the sensor-digitizer used. Use the spectral option in MULPLT to make a spectrum. When type of spectrum is asked, type 'n', this will now plot the power density acceleration spectrum which directly can be compared to the Anderson curve (Chapter 3).

- Make one plot for the broad band sensor and one for the SP (screen dumps), how do they compare to the Anderson curves, explain differences.

- Use one of the earlier 1 Hz measurements in nm and convert to noise power spectral values using formula 2 above assuming the filter to be 2/3 octave and compare to the curve made now.

Pensum

- All explanatory text in these experiment notes.

- General electronics. Ohms law, use of multimeter, internal resistance, power supply.  Pensum:.

- Seismic sensors: Mechanical and electrical constants, how to measure constants, overview of main types.

Knowledge of approximate values of generator constants and free period for different types of sensors.

- Signal conditioning: Amplifier general properties, filters.

- Analog recording: Typical units (heliocorder, MEQ800), typical gain and characteristics.

- Digital instruments: What is an A/D converter, what does number of bits mean, typical units on the market, anti alias filters,  SEISLOG station.

- Triggered systems: Detector and data storage.

- Seismic noise: Generally expected shape of noise curves, noise measured as displacement or power spectral density of acceleration.

Pensum from book:

Chapters 1, 2.1-2.5, 2.7-2.13, 2.18-2.21, 3, 4, 5.1-5.9, 6-6.1, 6.4-6.8, 7, 8.1-8.7, 8.12,  10-10.2.5, 10.3.4, 10.5-10.6, A1 – A3.1, A4 – A5.

From book,  no detailed derivations and no technical details are pensum.

# How to make reports

- Lab reports can be hand written or typed.

- The report should be self explanatory.

- All questions must be answered.

- All calculations must be shown so it is possible to understand what has been done.

-Observations and numbers must be given with an appropriate number of - digits in accordance with measurement accuracy. Give uncertainty if possible and appropriate.

- State which equipment has been used.

- Make a drawing of the experimental setup when needed.

- Include screen dumps when needed.

Example of part of a report

Experiment 1 – GEOF 275           Name:                 Date:

Purpose: Lean to use a multimeter and calculate internal resistance of a battery.

Experiment 1.1

The voltage of the supplied standard battery was 1.477 ± 0.002 V.

The power supply had two ranges specified, 0-15 V and 0-30 V. The measured ranges were:

| Range, V | Range measured with digital multimeter, V | Range measured with Built in analog multimeter, V | Difference in maximum voltage between digitial and analog measurements, V |
| --- | --- | --- | --- |
| 0-15 | 0-28.55 ±0.01 | 0 – 27.5 ±0.3 | 1.1 ± 0.3 |
| 0-30 | 0-36.12 ±0.01 | 0 – 35 ±1 | 1.1 ± 1 |

The output voltages were much higher than given in manual.

The accuracy of measurements with the analog meter is lower than the accuracy of the  measurements with the digital multimeter and the errors are in the order of 1 V.

The power supply has a current limiting capability. This means that the current can ………

In order to get an estimate of the accuracy of the multimeters, the same voltage was measured at the same time with 5 multimeters of the same type. The nominal voltage used was 10 V and the measured voltages were:

Multimeter 1: 	9.57 V

Multimeter 2:	9.56 V

Multimeter 3:	9.59 V

Multimeter 4: 	9.61 V

Multimeter 5:	9.59 V

Average:  9.59 V

Standard deviation: 0.02 V

For the 10 V range, the manual states that the accuracy is 0.5 % ±1 digit. This means that the error can be:

0.005 * 9.59 V + 0.01 V = 0.06 V

From the measurements above, it seems that the variation between readings is much less than 0.06 V. However, since the input voltage is not exactly known, there is a possibility of a fixed error, equal for all instruments (unlikely). Conclusion: Multimeters are within specifications for the 10 V DC range.

Experiment 1.4

An experiment was set up to calculate the internal resistance of a 1.5 V battery, see figure below.

The voltage of the battery before connecting it to the resistor was Vunloaded =1.477 V.

The measurements in the circuit were:

Loaded battery voltage        Vloaded = 1.372 V

Voltage over resistor           VR       = 1.351 V

Voltage  over multimeter    VM       =  0.020 V

Current	                       I          =  0.135 A

It is seen that the sum of the voltages over the resistor and the multimeter is equal to ( within measurement accuracy) the voltage over the battery as expected.

The internal resistance of the battery is:

(Vunloaded - Vloaded)/I = (1.477 V – 1.372 V)/0.135 A = 0.777 ohm

The internal resistance of the multimeter is:

VM/I = 0.020 V/0.135 A = 0.15 ohm

The internal resistance of the multimeter is significant compared to the internal resistance of the battery and it cannot be ignored as also seen by the 0.020 V drop.